## Sheet1
Annex 4 | Technical Submission
Equipment | Technical Requirement | Country of Origin | Compliance Yes/No (specify how you meet the requirement)
Electrical Safety analyser | a. Proprietary names, brands, catalogues, or reference numbers are designated to establish a minimum standard for the goods.
b. The Supplier may offer material or equipment of similar (or higher) characteristics, type, quality, appearance, finish, method of production and performance.
c. Detailed Operations and Maintenance Manuals in English language for each appropriate unit of the supplied Goods are to be provided.
d. Evidence should be provided indicating that the supplier has technical personnel that are capable to fulfill the scope of services, if applicable or needed. These include installation, commissioning, demonstration, on-site training of clinical staff, on-site training of bio-medical staff for first line maintenance, performance supervision of on-site assembly with report and start-up of the supplied Goods.
e. All equipment must be the latest versions of the proposed model.
f. Equipment requiring modification to operate on electrical voltage rated 240/415 Volt 50 Hz, IEC Type I Electrical Plug will not be accepted.
g. All software updates should be provided free of cost during Warranty.
h. All spare parts and consumables of equipment offered must be available for seven (7) years after commissioning of goods.
i. The conformity to the latest ISO Certification, CE Approved, EC Marked, IEC and US FDA Approved standards (including amendments) specified must be confirmed by a 3rd party agency accredited by Standards Australia, Standards New Zealand or by an independent international accreditation agency such as SGS, TUV, Det Norske Veritas. Comply with EN 60601-1 Safety Standard. A color copy of the confirmation of conformity must be submitted with the bid submission.
j. A manufacturer’s specification sheet and photo for each product must be submitted. All manufacturers must be ISO- 9001–certified
k. If the item uses consumables / reagents etc., a current price list of all consumables used by the item of equipment and an indication of how long the prices of consumables will be guaranteed. The supplier should also indicate the most cost-effective way of providing consumables / reagents.
Functional Requirement
· An automated portable device with available test standards selection to ANSI/AAMI ES-1, IEC62353, IEC60601-1, and AN/NZS 3551 that can perform the following primary electrical safety tests for medical equipment rated 230VAC at 50Hz:
o Line (mains) voltage: (180 V - 264 V ac rms)
o Mains and point to point voltage
o Ground wire (protective earth) resistance: 0 - 2Ω (±2 % of reading + 0.015 Ω)
o Insulation resistance: 0.5MΩ - 100MΩ, with source test voltage 500Vdc or 250Vdc
o Device current: (0 A - 20 A)
o Leakage current ranges: (0 µA - 199.9 μA), (200 µA to 1999 μA), (2 mA to 10 mA)
o Earth leakage test
o Lead - ground (patient) leakage tests
o Lead - lead (patient auxiliary) leakage tests
o Mains on Applied Parts leakage tests
o Enclosure leakage test
o Point to point leakage test
o ECG simulation - at least BPM: 30, 60, 120, 180, and 240
· Large, easy-to-read display with adjustable contrast
· Wireless communication & removable memory card for fast and convenient data storage and exchange
· Type I plug, 230VAC power outlet
· Handle and tilt stand
· Include all essential accessories such as test lead set, probe etc.
· Language available in English
· Two-year extended warranty inclusive of first year calibration
· Carrying Case
Electrical Safety analyser for Palau | Technical Requirement
a. Proprietary names, brands, catalogues, or reference numbers are designated to establish a minimum standard for the goods.
b. The Supplier may offer material or equipment of similar (or higher) characteristics, type, quality, appearance, finish, method of production and performance.
c. Detailed Operations and Maintenance Manuals in English language for each appropriate unit of the supplied Goods are to be provided.
d. Evidence should be provided indicating that the supplier has technical personnel that are capable to fulfill the scope of services, if applicable or needed. These include installation, commissioning, demonstration, on-site training of clinical staff, on-site training of bio-medical staff for first line maintenance, performance supervision of on-site assembly with report and start-up of the supplied Goods.
e. All equipment must be the latest versions of the proposed model.
f. Equipment requiring modification to operate on electrical voltage rated 240/415 Volt 50 Hz, IEC Type I Electrical Plug will not be accepted.
g. All software updates should be provided free of cost during Warranty.
h. All spare parts and consumables of equipment offered must be available for seven (7) years after commissioning of goods.
i. The conformity to the latest ISO Certification, CE Approved, EC Marked, IEC and US FDA Approved standards (including amendments) specified must be confirmed by a 3rd party agency accredited by Standards Australia, Standards New Zealand or by an independent international accreditation agency such as SGS, TUV, Det Norske Veritas. Comply with EN 60601-1 Safety Standard. A color copy of the confirmation of conformity must be submitted with the bid submission.
j. A manufacturer’s specification sheet and photo for each product must be submitted. All manufacturers must be ISO- 9001–certified
k. If the item uses consumables / reagents etc., a current price list of all consumables used by the item of equipment and an indication of how long the prices of consumables will be guaranteed. The supplier should also indicate the most cost-effective way of providing consumables / reagents.
Functional Requirement
· Multifaceted automated portable FDA compliant device that can perform the following primary electrical safety tests for medical equipment rated 110VAC at 60Hz:
o Line (mains) voltage: (90 V - 132 V ac rms), (180 V - 264 V ac rms)
o Mains and point to point voltage
o Ground wire (protective earth) resistance: 0 - 2Ω (±2 % of reading + 0.015 Ω)
o Insulation resistance: 0.5MΩ - 100MΩ, with source test voltage 500Vdc or 250Vdc
o Insulation resistance test: Mains- protective earth, AP - PE, Mains-NE, AP-NE
o Device current: (0 A - 20 A)
o Leakage current ranges: (0 µA - 199.9 μA), (200 µA to 1999 μA), (2 mA to 10 mA)
o Earth leakage test
o Lead - ground (patient) leakage tests
o Lead - lead (patient auxiliary) leakage tests
o Mains on Applied Parts leakage tests
o Chassis leakage test
o Point to point leakage test
o ECG simulation - at least BPM: 30, 60, 120, 180, and 240
· Large, easy-to-read display with adjustable contrast
· Wireless communication & removable memory card for fast and convenient data storage and exchange
· Automated test sequences with compliance to US electrical safety standards using on-board automation - ANSI/AAMI ES1:1993 & NFPA-99
· Include all essential accessories such as test lead set, probe etc.
· 110 VAC power outlet
· Handle and tilt stand
· Language available in English
· One year warranty and inclusive of first year calibration
· Carrying Case
Patient/Vital Signs Simulator | Technical Requirement
a. Proprietary names, brands, catalogues, or reference numbers are designated to establish a minimum standard for the goods.
b. The Supplier may offer material or equipment of similar (or higher) characteristics, type, quality, appearance, finish, method of production and performance.
c. Detailed Operations and Maintenance Manuals in English language for each appropriate unit of the supplied Goods are to be provided.
d. Evidence should be provided indicating that the supplier has technical personnel that are capable to fulfill the scope of services, if applicable or needed. These include installation, commissioning, demonstration, on-site training of clinical staff, on-site training of bio-medical staff for first line maintenance, performance supervision of on-site assembly with report and start-up of the supplied Goods.
e. All equipment must be the latest versions of the proposed model.
f. Equipment requiring modification to operate on electrical voltage rated 240/415 Volt 50 Hz, IEC Type I Electrical Plug will not be accepted.
g. All software updates should be provided free of cost during Warranty.
h. All spare parts and consumables of equipment offered must be available for seven (7) years after commissioning of goods.
i. The conformity to the latest ISO Certification, CE Approved, EC Marked, IEC and US FDA Approved standards (including amendments) specified must be confirmed by a 3rd party agency accredited by Standards Australia, Standards New Zealand or by an independent international accreditation agency such as SGS, TUV, Det Norske Veritas. Comply with EN 60601-1 Safety Standard. A color copy of the confirmation of conformity must be submitted with the bid submission.
j. A manufacturer’s specification sheet and photo for each product must be submitted. All manufacturers must be ISO- 9001–certified
k. If the item uses consumables / reagents etc., a current price list of all consumables used by the item of equipment and an indication of how long the prices of consumables will be guaranteed. The supplier should also indicate the most cost-effective way of providing consumables / reagents.
Functional Requirement
· One portable device that can simulate multiple vital signs tests such as electrocardiograph (ECG), fetal ECG, arrhythmia, respiration, temperature, non-invasive blood pressure, invasive-blood pressure, cardiac output simulator, SpO2 tester, and is the premier SpO2 simulator to test rainbow multi-wavelength waveforms.
· Include all essential accessories kit, cables, adapters, connectors for various patient monitor brands such as SPO2 probe, NIBP tubes, NIBP cables, cardiac output, temperature probes etc.
· Type I plug for 230 VAC power outlet
· Language available in English
· One year warranty and inclusive of first year calibration
For the Bidder: [insert name of the company]
Signature:
Name of the representative: [insert name of the representative]
Title: [insert Title of the representative]
Date: [Click or tap to enter a date]